## Slide 1
CONTINUOUS ATMOSPHERIC MONITORING AND AEROSOL TYPING USING CE376 DUAL-WAVELENGTH LIDAR
EXPLORE THE CLIMATE
(1) R&D Department, CIMEL Electronique, Paris, 75011, France (2) LOA—Laboratoire d’Optique Atmosphérique, University of Lille, 59000 Lille, France (3) Izaña Atmospheric Research Centre, S/C de Tenerife, 38001, Spain
(4) Prokhorov General Physics Institute, Moscow, Russia
Ioana E. Popovici (1,2), Maria F. Sanchez-Barrero (1,2), Yenny Gonzalez Ramos (1,3), Stéphane Victori (1), Qiaoyun Hu (2), Igor Veselovskii (1,4), William Boissière (2), Thierry Podvin (2) and Philippe Goloub (2)
XII Workshop Lidar Measurements in Latin America (WLMLA) 7-12 April 2024

## Slide 2
CIMEL COMPANY
UNE VUE DE PARIS, LE 27 JUILLET 2018
PH CREDITS: MUSTAFA YALCIN / ANADOLU AGENCY / AFP
Our main mission is to be the link between the scientific community and the industrial world
French manufacturer for more than 50 years of remote sensing solutions for monitoring the atmosphere and the environment
Atmospheric monitoring solutions:
LIDAR, photometer, infrared radiometer,
algorithms
AAMS
RESEARCH / CLIMATe SCIENCE
Air quality

## Slide 3: CIMEL – OUR PRODUCTS
CE318-T photometer – standard AERONET instrument (> 1500 in total in the world)
CE376 LiDAR – compact automatic low power lidar (ALC)
AAMS / iAAMS – synergy and software
CE312 radiometer – 4 to 6 thermal IR channels for brightness temperature measurements
CE710 LiDAR – Raman fluorescence high-power LiDAR (AHL)

## Slide 4
CIMEL – PARTNERS ecosystem
Innovative LIDAR research
(O3, WV, HSRL)
Data inversion
Integration in iAAMS
CARS (Center for Aerosol Remote Sensing)
Calibration / SOP
Master calibration
WMO standardization
AERONET
Network instrument
Photometer evolutions
Optical technology (non-linear)
Industrial applications
Instrumental synergies
Mobile observations
Scientific dissemination
https://www.cimel.fr/
4

## Slide 5: CE376 DUAL-WAVELENGTH DEPOLARIZATION LIDAR
Size: 71 x 46 x 69 cm
Weight: 35 kg
Continuous, 24/7
Autonomous, turn-key
Eye-safe, IEC 60825-1 compliant
No consummables needed
532 nm Nd:YAG laser
29 mW, 6 µJ, 4.7 kHz, 0.2 nm
808 nm diode laser
20 mW, 4 µJ, 4.7 kHz, 0.6 nm
Depolarization: 532 nm
FOV (E/R): 420/550 µrad
One-block E/R modules
Detection by APD: 532 nm parallel/cross, 808 nm
Full overlap: 500 m
Altitude range (AOD = 0.3)
at 532 nm
night: 20 km / day: 6 km
at 808 nm
night: 10 km / day: 4 km

## Slide 6: CE376 – CONTINUOUS OBSERVATIONS
Continuous lidar measurements are valuable to detect aerosol events above one station

## Slide 7: CE376 – CONTINUOUS OBSERVATIONS
Continuous lidar measurements are valuable to detect aerosol events above one station
Depolarization and 2-wavelengths help identify aerosol types (dust, smoke, clouds etc.)
dust
rain
ice clouds / dust
pollution

## Slide 8: DUST EVENT – LILLE, FRANCE
Saharan dust transport (1-5 km altitude) at ATOLL ACTRIS station in Lille, France, 31 March-2 April 2021
AOD532 nm ↗ 1 , Angstrom Exponent ↘︎ 0.2 (coarse)
Comparison with LILAS, reference ACTRIS Raman multi-wavelength lidar
Sanchez Barrero et al., AMT, 2024

## Slide 9: DUST EVENT – LILLE, FRANCE
Sanchez Barrero et al., AMT, 2024
Extinction (⍺) / backscatter (β) profiles comparable to LILAS; differences due to constant LR (Klett)
Low spectral dependence (⍺ and β) - dust
VDR =PLDR = 0.36 ± 0.05 (high depol.)
EAE (Ang. Exp.) = 0.4 ± 0.1 (coarse), AEphot=0.2-0.7
ACR = 0.42 ± 0.05, CR = 0.69 ± 0.14 (coarse)
LR532 nm= 54 sr, LR808 nm= 69 sr
Notes: LR of LILAS: 48 sr at 1-3 km and 72 sr at 3.3-4.7 km
ACR and CR: 0.42 ± 0.05 and 0.69 ± 0.14, for the lower layer and 0.38 ± 0.04 and 0.65 ± 0.12 for the second layer
METIS showed VLDR values 10 % higher than LILAS

## Slide 10: VOLCANIC EVENT – IZANA, SPAIN
Cumbre Vieja eruption on La Palma Island (21 September 2021) detected at Izaña Observatory, Spain (2373 m AGL, in the FT)
High concentrations of , and observed at ground level in La Palma and Izaña
Continuous CE376 lidar allowed detection of volcanic plumes arrival at 3-5 km ASL
140 km
Sanchez Barrero, thesis, 2024
Notes: high emissions of reaching up to 20 DU (TROPOMI) spectrometer on board Sentinel-5 satellite
high emissions of 𝑆𝑂_2 reaching up to 20 DU (TROPOMI) spectrometer on board Sentinel-5 satellite

## Slide 11: VOLCANIC EVENT – IZANA, SPAIN
AOD532 nm ↗ 0.17 , Angs. Exp. ↗ 2 (fine)
High spectral dependence (⍺ and β)
at 532 nm / at 808 nm
VDR =PLDR = 0.03 (low depol.)
EAE (Ang. Exp.) = 1.8 (fine), AEphot=1.8
ACR = 0.33, CR = 0.33 (fine)
LR532 nm= 38 sr, LR808 nm= 59 sr
Sanchez Barrero, thesis, 2024
Non-ash
(sulfates)
Notes: AOD532 nm= 0.17, AOD808 nm= 0.08

## Slide 12: DUST / SMOKE EVENT – LILLE, FRANCE
Complex case of Saharan dust mixed with smoke from wildfires in Portugal, Spain and SW France (Gironde) observed at ATOLL, Lille on 17-19 July 2022
Dust Optical Depth (DOD) of 0.1-0.2 forecasted over France
HYSPLIT back-trajectories for the layers confirm that different air masses are impacting the site
Sanchez Barrero, AGU Fall Meeting, 2022
Notes: Lidar observations present complex vertical distribution with variable mixing of dust and smoke, showing us the importance to have access to information from both instruments
WMO Barcelona Dust Regional Center
Back-trajectories ending at 3 altitudes, 1.7 km (L1 in red), 2.7 km (L2 in blue) and 4 km (L3 in green)

## Slide 13: DUST / SMOKE EVENT – LILLE, FRANCE
Sanchez Barrero, AGU Fall Meeting, 2022
Dust
Smoke
Dust /Smoke
AOD532 nm = 0.3 , Angs. Exp. = 1 (mixed)
Low PLDR () for L2 – transported smoke
High PLDR for L1 () and L3 () - dust mixed with smoke
EAE (Ang. Exp.) = 0.83 (mixed), AEphot=0.9
ACR = 0.54, CR = 0.74 for L3 (coarse)
LR532 nm= 39 sr, LR808 nm= 38 sr
L1
L2
L3
Notes: L1 (1.5-2.2 km asl), L2 (2.4-3.2 km asl) and L3 (3.2-4.5 km asl).

## Slide 14: CE376 - OBS4CLIM PROJECT
Extend the spatial and temporal coverage of observations and complete the research platforms (in Southern Hemisphere, remote/challenging environments)
Consolidate the capacity of atmospheric RIs to provide data and products responding to the needs of users and future challenges of global observation
GP
GPN
GPN
https://www.obs4clim.fr/
GP
GP = Green Polarized
GPN = Green Polarized NIR
4 x CE376 – summer 2024
1 x CE710 – summer 2024
CE710
Notes: Prepare the research infrastructures (RIs) to respond to future challenges of global observation
Strengthen the attractiveness of atmospheric research platforms and enhance access

## Slide 15: CE710 MULTI-WAVELENGTH RAMAN FLUORESCENCE LIDAR
Size: 2 x 1.2 x 1.1 m
Weight: >250 kg
Scheduled measurements
Need of roof hatch
Rain sensor shutdown
Automatic calibrator
Mie: 355, 532, 1064
Raman: 353, 387, 530, 1056
Depolarization: 355, 532, 1064
Water Vapor: 408
Fluorescence: 420-640 nm
Configurable «LEGO-like» analyzer (2, 5, 15 channels etc.)
Follows ACTRIS guidelines
Laser: Lumibird (flash lamp, DPSS) - configurable
LIFE - Merion MW 9-100, DPSS laser, 200 mJ @ 355 nm, 100 Hz
Telescope: Newton, 40 cm
Adjustable FOV: pinhole
Built with I. Veselovskyi

## Slide 16: CE710 – REFINE AEROSOL TYPING
LILAS backscatter and fluorescence observations at ATOLL, Lille, France on 9 June 2023
Better discrimination of organic particles (smoke, pollution, pollen)
Backscatter (532 nm)
Fluorescence (460 nm)
Water Vapor (408 nm)
Depolarization (532 nm)
(See Igor Veseloskyi’s talks)
AUSTRAL software (Hu, 2024)

## Slide 17: CONCLUSIONS
Saharan dust transport at Lille, France, characterized using depolarization and 2-wl CE376 LiDAR (high PLDR, low EAE)
Volcanic emission aerosols (non-ash/sulfates) from Cumbre Vieja observed at Izaña, Spain, using CE376 automatic LiDAR (low PLDR, high EAE)
Advantage of continuous LiDAR measurements for detection of aerosol events
Validation of CE376 LiDAR against reference Raman lidar in ACTRIS (LILAS, Lille, France)
Difficulties in retrieving aerosol properties in complex cases using Klett method (mixed dust/smoke case, limit of vertical constant LR) – limitation of CE376 elastic LiDAR
Multi-wavelength Raman approach, depolarization and fluorescence channels help better discriminate aerosol types vertically (LILAS, example of CE710 LiDAR)
Manufacturing of state-of-the-art high power LiDAR (CE710) in collaboration with renowned experts

## Slide 18
THANK YOU
21/03/2019
i-popovici@cimel.fr
https://www.cimel.fr/